## summary
data_type | gantry | gantry_angle | energy | fwhm
tps | 1 | 90 | 70 | 13.5625 | Sheet | Checked | I have checked each sheet's data with comparison to the source data (location at the top of each individual sheet). This includes TPS data, plus gantry specific spot data for each of the four gantries. To the best of my ability, I am happy that they are an accurate representation of the data, save for some differences in rounding. HC, 01/06/2026.
tps | 1 | 90 | 100 | 11.1772 | TPS | x
tps | 1 | 90 | 150 | 9.0818 | G1_GA0 | x
tps | 1 | 90 | 200 | 8.1719 | G1_GA90 | x
tps | 1 | 90 | 240 | 7.4099 | G1_GA180 | x
baseline | 1 | 0 | 70 | 13.6962 | G1_GA270 | x
baseline | 1 | 0 | 100 | 11.2674 | G2_GA0 | x
baseline | 1 | 0 | 150 | 9.2442 | G2_GA0_annual | x
baseline | 1 | 0 | 200 | 8.2459 | G2_GA90 | x
baseline | 1 | 0 | 240 | 7.7621 | G2_GA180 | x
baseline | 1 | 90 | 70 | 13.904 | G2_GA270 | x
baseline | 1 | 90 | 100 | 11.1578 | G3_GA0 | x
baseline | 1 | 90 | 150 | 9.3372 | G3_GA90 | x
baseline | 1 | 90 | 200 | 8.3989 | G3_GA180 | x
baseline | 1 | 90 | 240 | 7.78 | G3_GA270 | x
baseline | 1 | 180 | 70 | 13.9158 | G4_GA0 | x
baseline | 1 | 180 | 100 | 11.3669 | G4_GA90 | x
baseline | 1 | 180 | 150 | 9.309 | G4_GA180 | x
baseline | 1 | 180 | 200 | 8.5053 | G4_GA270 | x
baseline | 1 | 180 | 240 | 7.9821
baseline | 1 | 270 | 70 | 13.9654
baseline | 1 | 270 | 100 | 11.2928
baseline | 1 | 270 | 150 | 9.36
baseline | 1 | 270 | 200 | 8.5463
baseline | 1 | 270 | 240 | 7.9485
baseline | 2 | 0 | 70 | 14.2295
baseline | 2 | 0 | 100 | 11.4674
baseline | 2 | 0 | 150 | 9.4682
baseline | 2 | 0 | 200 | 8.742
baseline | 2 | 0 | 240 | 8.0884
baseline | 2 | 90 | 70 | 14.272
baseline | 2 | 90 | 100 | 11.4813
baseline | 2 | 90 | 150 | 9.5962
baseline | 2 | 90 | 200 | 8.7733
baseline | 2 | 90 | 240 | 7.2871
baseline | 2 | 180 | 70 | 14.4119
baseline | 2 | 180 | 100 | 11.5536
baseline | 2 | 180 | 150 | 9.4759
baseline | 2 | 180 | 200 | 8.6086
baseline | 2 | 180 | 240 | 7.9953
baseline | 2 | 270 | 70 | 14.3666
baseline | 2 | 270 | 100 | 11.6638
baseline | 2 | 270 | 150 | 9.5685
baseline | 2 | 270 | 200 | 8.6907
baseline | 2 | 270 | 240 | 7.2837
baseline | 3 | 0 | 70 | 14.083
baseline | 3 | 0 | 100 | 11.5599
baseline | 3 | 0 | 150 | 9.4116
baseline | 3 | 0 | 200 | 8.2472
baseline | 3 | 0 | 240 | 7.5387
baseline | 3 | 90 | 70 | 14.0151
baseline | 3 | 90 | 100 | 11.727
baseline | 3 | 90 | 150 | 9.4566
baseline | 3 | 90 | 200 | 8.41
baseline | 3 | 90 | 240 | 7.8114
baseline | 3 | 180 | 70 | 14.0475
baseline | 3 | 180 | 100 | 11.7108
baseline | 3 | 180 | 150 | 9.3601
baseline | 3 | 180 | 200 | 8.1876
baseline | 3 | 180 | 240 | 7.4849
baseline | 3 | 270 | 70 | 13.9989
baseline | 3 | 270 | 100 | 11.7629
baseline | 3 | 270 | 150 | 9.4952
baseline | 3 | 270 | 200 | 8.6767
baseline | 3 | 270 | 240 | 7.8871
baseline | 4 | 0 | 70 | 13.7825
baseline | 4 | 0 | 100 | 11.3092
baseline | 4 | 0 | 150 | 9.3456
baseline | 4 | 0 | 200 | 8.2749
baseline | 4 | 0 | 240 | 7.5158
baseline | 4 | 90 | 70 | 13.8277
baseline | 4 | 90 | 100 | 11.396
baseline | 4 | 90 | 150 | 9.3452
baseline | 4 | 90 | 200 | 8.1836
baseline | 4 | 90 | 240 | 7.7188
baseline | 4 | 180 | 70 | 13.8377
baseline | 4 | 180 | 100 | 11.338
baseline | 4 | 180 | 150 | 9.3401
baseline | 4 | 180 | 200 | 8.3253
baseline | 4 | 180 | 240 | 7.4852
baseline | 4 | 270 | 70 | 13.625
baseline | 4 | 270 | 100 | 11.3255
baseline | 4 | 270 | 150 | 9.4782
baseline | 4 | 270 | 200 | 8.649
baseline | 4 | 270 | 240 | 7.5844

## TPS
TPS spot size without RS
C:\Users\KALONGMA\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\Treatment_Planning\Gantry1\Profiles\Fits_G90_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.028288 | 0.3229 | 5.8386 | 5.704 | 5.7713 | 13.5625 | 0 | 0.91219
75 | 1 | 0.21298 | 0.30028 | 5.6284 | 5.4421 | 5.5353 | 13.0079 | 0 | 0.98422
80 | 1 | 0.050231 | 0.55018 | 5.4605 | 5.2555 | 5.358 | 12.5913 | 0 | 0.99441
85 | 1 | 0.094419 | 0.3201 | 5.1857 | 5.0012 | 5.0935 | 11.9697 | 0 | 0.98649
90 | 1 | 0.10433 | 0.15915 | 5.0863 | 4.9092 | 4.9978 | 11.7448 | 0 | 0.99571
95 | 0.9944 | 0.2784 | 0.23071 | 4.9339 | 4.7752 | 4.8545 | 11.4082 | 0 | 0.99544
100 | 1 | 0.3841 | 0.40817 | 4.8301 | 4.6824 | 4.7563 | 11.1772 | 0 | 0.86258
105 | 1 | 0.18108 | 0.40348 | 4.702 | 4.5795 | 4.6408 | 10.9058 | 0 | 0.88678
110 | 1 | 0.26715 | 0.28289 | 4.5743 | 4.4813 | 4.5278 | 10.6403 | 0 | 0.95667
115 | 1 | 0.13412 | 0.11757 | 4.4597 | 4.3739 | 4.4168 | 10.3794 | 0 | 0.84675
120 | 1 | 0.35187 | 0.27326 | 4.403 | 4.3499 | 4.3764 | 10.2846 | 0 | 0.90516
125 | 1 | 0.18764 | 0.35083 | 4.2734 | 4.257 | 4.2652 | 10.0232 | 0 | 0.88132
130 | 1 | 0.3329 | 0.4645 | 4.1941 | 4.2133 | 4.2037 | 9.8788 | 0 | 0.93839
135 | 1 | 0.030465 | 0.31399 | 4.0724 | 4.1187 | 4.0955 | 9.6245 | 0 | 0.92365
140 | 1 | -0.013047 | 0.26854 | 4.0341 | 4.1001 | 4.0671 | 9.5577 | 0 | 0.94314
145 | 1 | 0.28487 | 0.26268 | 3.9242 | 4.0185 | 3.9713 | 9.3327 | 0 | 0.8646
150 | 1 | 0.084259 | 0.46811 | 3.8091 | 3.9202 | 3.8646 | 9.0818 | 0 | 0.89737
155 | 1 | 0.15035 | 0.39528 | 3.7679 | 3.8998 | 3.8339 | 9.0096 | 0 | 1
160 | 1 | -0.10135 | 0.50357 | 3.7138 | 3.8542 | 3.784 | 8.8924 | 0 | 0.99916
165 | 1 | 0.16675 | 0.41271 | 3.6523 | 3.7986 | 3.7254 | 8.7548 | 0 | 0.99896
170 | 1 | 0.19332 | 0.40284 | 3.6545 | 3.7933 | 3.7239 | 8.7513 | 0 | 0.99948
175 | 1 | 0.064324 | 0.153 | 3.63 | 3.763 | 3.6965 | 8.6868 | 0 | 1
180 | 1 | -0.056354 | 0.18818 | 3.5846 | 3.6988 | 3.6417 | 8.558 | 0 | 0.99988
185 | 1 | 0.011874 | 0.4186 | 3.5691 | 3.6351 | 3.6021 | 8.465 | 0 | 1
190 | 1 | -0.046579 | 0.454 | 3.5534 | 3.5631 | 3.5582 | 8.3618 | 0 | 1
195 | 1 | -0.10572 | 0.30437 | 3.5131 | 3.4808 | 3.4969 | 8.2178 | 0 | 1
200 | 1 | 0.16017 | 0.34579 | 3.5093 | 3.4455 | 3.4774 | 8.1719 | 0 | 1
205 | 1 | -0.022991 | 0.14955 | 3.4682 | 3.388 | 3.4281 | 8.0561 | 0 | 1
210 | 1 | -0.011828 | 0.25474 | 3.4327 | 3.3545 | 3.3936 | 7.9749 | 0 | 1
215 | 1 | -0.071066 | 0.34532 | 3.384 | 3.3164 | 3.3502 | 7.873 | 0 | 0.99954
220 | 1 | -0.074268 | 0.46978 | 3.3979 | 3.3388 | 3.3684 | 7.9157 | 0 | 1
225 | 1 | 0.072405 | 0.28963 | 3.3376 | 3.293 | 3.3153 | 7.791 | 0 | 1
230 | 1 | -0.080196 | 0.40967 | 3.3245 | 3.2925 | 3.3085 | 7.775 | 0 | 1
235 | 1 | 0.00089833 | 0.14072 | 3.2791 | 3.2543 | 3.2667 | 7.6768 | 0 | 1
240 | 1 | -0.069989 | 0.29497 | 3.1615 | 3.1448 | 3.1531 | 7.4099 | 0 | 1
245 | 1 | -0.0078748 | 0.30069 | 3.0545 | 2.9046 | 2.9795 | 7.0019 | 0 | 1

## G1_GA0
Gantry 1 GA0 data
T:\Routine QA\Annual Checks\G1\2023\Profiles\Fits_G0_RS0_Dist0_2023.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.099645 | 0.67425 | 5.7891 | 5.8672 | 5.8282 | 13.6962 | 0 | 0.7
75 | 1 | 0.069502 | 0.341 | 5.5187 | 5.5982 | 5.5584 | 13.0623 | 0 | 0.95255
80 | 1 | 0.12984 | 0.41343 | 5.3617 | 5.4257 | 5.3937 | 12.6752 | 0 | 0.99044
85 | 1 | 0.31515 | 0.39252 | 5.1823 | 5.228 | 5.2051 | 12.2321 | 0 | 0.99197
90 | 1 | 0.26923 | 0.30933 | 5.0835 | 5.0994 | 5.0914 | 11.9649 | 0 | 0.86664
95 | 1 | 0.078897 | 0.12375 | 4.9158 | 4.9129 | 4.9144 | 11.5487 | 0 | 0.99563
100 | 1 | 0.34723 | 0.30153 | 4.8026 | 4.7866 | 4.7946 | 11.2674 | 0 | 0.85166
105 | 1 | 0.09109 | 0.61336 | 4.6466 | 4.6207 | 4.6337 | 10.8891 | 0 | 0.8727
110 | 1 | 0.42226 | 0.39032 | 4.601 | 4.5727 | 4.5868 | 10.779 | 0 | 0.86479
115 | 1 | 0.1055 | 0.26421 | 4.4625 | 4.4266 | 4.4445 | 10.4447 | 0 | 0.87171
120 | 1 | 0.38229 | 0.35201 | 4.3476 | 4.2993 | 4.3234 | 10.1601 | 0 | 0.83377
125 | 1 | 0.44307 | 0.17445 | 4.3243 | 4.2648 | 4.2946 | 10.0923 | 0 | 0.86216
130 | 1 | 0.1749 | 0.72838 | 4.2657 | 4.1891 | 4.2274 | 9.9344 | 0 | 0.93522
135 | 1 | 0.23663 | 0.16781 | 4.1961 | 4.1068 | 4.1515 | 9.756 | 0 | 0.87504
140 | 1 | 0.020382 | 0.050146 | 4.0921 | 3.9819 | 4.037 | 9.487 | 0 | 1
145 | 1 | 0.1612 | 0.06558 | 4.04 | 3.9134 | 3.9767 | 9.3453 | 0 | 0.99942
150 | 1 | 0.16782 | 0.17395 | 4.0069 | 3.8605 | 3.9337 | 9.2442 | 0 | 1
155 | 1 | -0.0036872 | 0.538 | 3.9741 | 3.8037 | 3.8889 | 9.1388 | 0 | 0.97706
160 | 1 | 0.19997 | 0.19447 | 3.9583 | 3.7709 | 3.8646 | 9.0818 | 0 | 1
165 | 1 | 0.12058 | 0.36721 | 3.8766 | 3.6907 | 3.7836 | 8.8915 | 0 | 0.92473
170 | 1 | 0.25843 | 0.44747 | 3.8169 | 3.6369 | 3.7269 | 8.7582 | 0 | 1
175 | 1 | 0.1285 | 0.19503 | 3.7588 | 3.5934 | 3.6761 | 8.6388 | 0 | 1
180 | 1 | 0.24318 | 0.4652 | 3.712 | 3.564 | 3.638 | 8.5492 | 0 | 1
185 | 1 | 0.097075 | 0.23683 | 3.7072 | 3.5789 | 3.6431 | 8.5613 | 0 | 0.94208
190 | 1 | 0.10845 | 0.28617 | 3.6538 | 3.5503 | 3.602 | 8.4647 | 0 | 1
195 | 1 | 0.20593 | 0.41299 | 3.5959 | 3.519 | 3.5574 | 8.3599 | 0 | 1
200 | 1 | 0.37566 | 0.13345 | 3.542 | 3.4758 | 3.5089 | 8.2459 | 0 | 1
205 | 1 | 0.17579 | 0.38333 | 3.5451 | 3.4911 | 3.5181 | 8.2675 | 0 | 1
210 | 1 | 0.30479 | 0.32732 | 3.5169 | 3.4492 | 3.4831 | 8.1852 | 0 | 1
215 | 1 | 0.18253 | 0.36432 | 3.5159 | 3.4387 | 3.4773 | 8.1716 | 0 | 1
220 | 1 | 0.12172 | 0.29092 | 3.4527 | 3.3674 | 3.4101 | 8.0137 | 0 | 1
225 | 1 | 0.27807 | 0.085703 | 3.461 | 3.3559 | 3.4084 | 8.0098 | 0 | 1
230 | 1 | 0.27514 | 0.25891 | 3.4412 | 3.3176 | 3.3794 | 7.9415 | 0 | 1
235 | 1 | 0.082368 | 0.27266 | 3.4178 | 3.2712 | 3.3445 | 7.8596 | 0 | 1
240 | 1 | 0.34669 | 0.12148 | 3.3921 | 3.2139 | 3.303 | 7.7621 | 0 | 1
245 | 1 | -0.10436 | -0.22547 | 3.2261 | 3.1495 | 3.1878 | 7.4913 | 0 | 1

## G1_GA90
Gantry 1 Gantry angle 90
T:\Routine QA\Annual Checks\G1\2023\Profiles\GA90\Fits_G90_RS0_Dist0_2023.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.37344 | 0.3635 | 5.9861 | 5.847 | 5.9166 | 13.904 | 0 | 0.98811
75 | 1 | 0.28226 | 0.52938 | 5.7064 | 5.5207 | 5.6136 | 13.1919 | 0 | 0.99438
80 | 1 | 0.059787 | 0.45642 | 5.495 | 5.2965 | 5.3957 | 12.68 | 0 | 0.88984
85 | 1 | 0.082372 | 0.027352 | 5.2426 | 5.0472 | 5.1449 | 12.0905 | 0 | 0.89287
90 | 1 | 0.082781 | 0.53339 | 5.1495 | 4.9603 | 5.0549 | 11.8791 | 0 | 0.88525
95 | 1 | 0.28075 | 0.10721 | 5.0271 | 4.8657 | 4.9464 | 11.624 | 0 | 1
100 | 1 | 0.016068 | -0.041157 | 4.8189 | 4.6771 | 4.748 | 11.1578 | 0 | 0.90682
105 | 1 | 0.068781 | 0.33713 | 4.7526 | 4.6294 | 4.691 | 11.0238 | 0 | 0.91125
110 | 1 | 0.28302 | 0.34068 | 4.6356 | 4.5312 | 4.5834 | 10.771 | 0 | 0.88511
115 | 1 | 0.38879 | 0.51964 | 4.483 | 4.3938 | 4.4384 | 10.4302 | 0 | 0.8836
120 | 1 | 0.3188 | 0.25528 | 4.4236 | 4.3634 | 4.3935 | 10.3247 | 0 | 0.90426
125 | 1 | -0.021555 | -0.10358 | 4.3013 | 4.2733 | 4.2873 | 10.0752 | 0 | 0.90044
130 | 1 | 0.19187 | -0.047726 | 4.223 | 4.2388 | 4.2309 | 9.9426 | 0 | 0.89658
135 | 1 | 0.20217 | 0.2893 | 4.1837 | 4.228 | 4.2059 | 9.8838 | 0 | 0.8762
140 | 1 | 0.28925 | 0.22169 | 4.0816 | 4.1594 | 4.1205 | 9.6832 | 0 | 0.92194
145 | 1 | 0.25379 | 0.20176 | 3.9343 | 4.045 | 3.9897 | 9.3757 | 0 | 0.88089
150 | 1 | 0.086478 | 0.1036 | 3.9069 | 4.0396 | 3.9733 | 9.3372 | 0 | 1
155 | 1 | 0.062562 | 0.078098 | 3.8185 | 3.9742 | 3.8964 | 9.1565 | 0 | 1
160 | 1 | 0.31101 | 0.0016956 | 3.817 | 3.9759 | 3.8964 | 9.1566 | 0 | 0.92018
165 | 1 | 0.3146 | 0.073615 | 3.7853 | 3.9438 | 3.8645 | 9.0816 | 0 | 1
170 | 1 | 0.00098335 | 0.097591 | 3.7446 | 3.8895 | 3.8171 | 8.9701 | 0 | 1
175 | 1 | -0.07391 | 0.10107 | 3.7348 | 3.8748 | 3.8048 | 8.9413 | 0 | 1
180 | 1 | 0.18676 | 0.3272 | 3.6985 | 3.8136 | 3.7561 | 8.8268 | 0 | 0.94513
185 | 1 | 0.18433 | 0.064551 | 3.6762 | 3.7477 | 3.7119 | 8.723 | 0 | 0.93167
190 | 1 | -0.040648 | -0.08068 | 3.6368 | 3.6513 | 3.644 | 8.5634 | 0 | 1
195 | 1 | 0.18381 | 0.23018 | 3.6582 | 3.621 | 3.6396 | 8.553 | 0 | 0.95949
200 | 1 | 0.10519 | 0.16156 | 3.6083 | 3.5396 | 3.574 | 8.3989 | 0 | 0.9424
205 | 1 | 0.19715 | -0.14688 | 3.5928 | 3.5156 | 3.5542 | 8.3524 | 0 | 0.94657
210 | 1 | 0.26386 | 0.12096 | 3.5669 | 3.4856 | 3.5263 | 8.2867 | 0 | 1
215 | 1 | 0.25937 | -0.095398 | 3.526 | 3.4674 | 3.4967 | 8.2173 | 0 | 1
220 | 1 | 0.029445 | -0.089819 | 3.5028 | 3.4638 | 3.4833 | 8.1858 | 0 | 1
225 | 1 | 0.26313 | -0.012009 | 3.4786 | 3.4566 | 3.4676 | 8.1489 | 0 | 1
230 | 1 | 0.061119 | 0.10611 | 3.4522 | 3.4464 | 3.4493 | 8.1059 | 0 | 1
235 | 1 | 0.093464 | 0.12261 | 3.3901 | 3.3877 | 3.3889 | 7.9639 | 0 | 1
240 | 1 | 0.1513 | -0.15666 | 3.3076 | 3.3137 | 3.3107 | 7.78 | 0 | 1
245 | 1 | 0.17789 | 0.051054 | 3.1521 | 3.0011 | 3.0766 | 7.23 | 0 | 1

## G1_GA180
Gantry 1 Gantry angle 180
T:\Routine QA\Annual Checks\G1\2023\Profiles\GA180\Fits_G180_RS0_Dist0_2023.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.16689 | 0.64914 | 5.9489 | 5.8944 | 5.9216 | 13.9158 | 0 | 0.98619
75 | 1 | 0.18347 | 0.35156 | 5.6576 | 5.6337 | 5.6456 | 13.2673 | 0 | 0.91521
80 | 1 | 0.15736 | 0.29541 | 5.4019 | 5.3842 | 5.393 | 12.6736 | 0 | 0.88919
85 | 1 | -0.045031 | 0.41811 | 5.2194 | 5.1961 | 5.2078 | 12.2383 | 0 | 0.88474
90 | 1 | 0.26 | 0.13207 | 5.1168 | 5.0892 | 5.103 | 11.9921 | 0 | 0.88255
95 | 1 | 0.23738 | 0.55017 | 4.9534 | 4.9143 | 4.9339 | 11.5946 | 0 | 0.83711
100 | 1 | 0.20375 | 0.33208 | 4.8639 | 4.81 | 4.837 | 11.3669 | 0 | 0.84366
105 | 1 | -0.069472 | 0.036826 | 4.7653 | 4.7099 | 4.7376 | 11.1334 | 0 | 0.86475
110 | 1 | 0.14799 | 0.35652 | 4.5719 | 4.5079 | 4.5399 | 10.6688 | 0 | 0.89553
115 | 1 | 0.19493 | 0.23119 | 4.5638 | 4.4928 | 4.5283 | 10.6415 | 0 | 0.90872
120 | 1 | 0.17647 | 0.23647 | 4.4687 | 4.3847 | 4.4267 | 10.4026 | 0 | 0.87629
125 | 1 | -0.14777 | 0.091964 | 4.3714 | 4.2679 | 4.3197 | 10.1512 | 0 | 0.91241
130 | 1 | -0.1359 | 0.40999 | 4.3229 | 4.1991 | 4.261 | 10.0133 | 0 | 0.90347
135 | 1 | 0.14698 | 0.35814 | 4.2124 | 4.0723 | 4.1423 | 9.7344 | 0 | 0.90244
140 | 1 | 0.049116 | 0.33702 | 4.2058 | 4.0464 | 4.1261 | 9.6964 | 0 | 0.87642
145 | 1 | -0.054646 | 0.11985 | 4.1325 | 3.9636 | 4.0481 | 9.513 | 0 | 0.8957
150 | 1 | -0.062552 | 0.43609 | 4.0549 | 3.8677 | 3.9613 | 9.309 | 0 | 0.91138
155 | 1 | -0.078098 | 0.38015 | 4.0292 | 3.8208 | 3.925 | 9.2237 | 0 | 0.92345
160 | 1 | -0.043284 | 0.12676 | 3.994 | 3.7736 | 3.8838 | 9.127 | 0 | 1
165 | 1 | -0.061333 | 0.21364 | 3.9411 | 3.7293 | 3.8352 | 9.0127 | 0 | 1
170 | 1 | 0.20318 | 0.39723 | 3.8702 | 3.6657 | 3.7679 | 8.8546 | 0 | 0.95638
175 | 1 | 0.18115 | 0.14947 | 3.8555 | 3.6583 | 3.7569 | 8.8286 | 0 | 1
180 | 1 | -0.078456 | 0.1295 | 3.8331 | 3.6515 | 3.7423 | 8.7944 | 0 | 0.9158
185 | 1 | 0.098871 | 0.075208 | 3.7555 | 3.6028 | 3.6792 | 8.646 | 0 | 1
190 | 1 | -0.13557 | 0.14691 | 3.7187 | 3.5965 | 3.6576 | 8.5953 | 0 | 1
195 | 1 | -0.0093257 | 0.22995 | 3.667 | 3.5721 | 3.6195 | 8.5059 | 0 | 0.93977
200 | 1 | 0.072581 | 0.022278 | 3.6621 | 3.5764 | 3.6193 | 8.5053 | 0 | 0.95378
205 | 1 | -0.14609 | -0.012797 | 3.6343 | 3.5473 | 3.5908 | 8.4385 | 0 | 1
210 | 1 | -0.12766 | 0.16303 | 3.6273 | 3.533 | 3.5802 | 8.4134 | 0 | 0.94723
215 | 1 | 0.050263 | -0.11938 | 3.5848 | 3.4815 | 3.5331 | 8.3029 | 0 | 1
220 | 1 | 0.13201 | 0.009761 | 3.5404 | 3.4353 | 3.4878 | 8.1964 | 0 | 1
225 | 1 | -0.010605 | 0.22371 | 3.568 | 3.4591 | 3.5135 | 8.2568 | 0 | 0.98612
230 | 1 | -0.05973 | -0.19853 | 3.5162 | 3.3997 | 3.458 | 8.1262 | 0 | 0.97898
235 | 1 | -0.022124 | -0.06218 | 3.4762 | 3.3518 | 3.414 | 8.0229 | 0 | 0.98768
240 | 1 | 0.21538 | -0.16831 | 3.4692 | 3.3241 | 3.3967 | 7.9821 | 0 | 1
245 | 1 | 0.080157 | -0.077908 | 3.2988 | 3.2532 | 3.276 | 7.6985 | 0 | 1

## G1_GA270
Gantry 1 GA270
T:\Routine QA\Annual Checks\G1\2023\Profiles\GA270\Fits_G270_RS0_Dist0_2023.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.40409 | 0.65826 | 5.8928 | 5.9926 | 5.9427 | 13.9654 | 0 | 0.99443
75 | 1 | 0.1465 | 0.45131 | 5.6088 | 5.7151 | 5.6619 | 13.3056 | 0 | 0.9122
80 | 1 | -0.066347 | -0.001207 | 5.3613 | 5.4794 | 5.4204 | 12.7378 | 0 | 0.92184
85 | 1 | 0.1042 | 0.18504 | 5.1174 | 5.2332 | 5.1753 | 12.1619 | 0 | 0.89653
90 | 1 | 0.45286 | 0.4913 | 4.9756 | 5.096 | 5.0358 | 11.8342 | 0 | 0.8881
95 | 1 | 0.17751 | 0.31567 | 4.8834 | 5.0208 | 4.9521 | 11.6374 | 0 | 0.84266
100 | 1 | 0.3848 | 0.61513 | 4.7344 | 4.8766 | 4.8055 | 11.2928 | 0 | 0.86358
105 | 1 | 0.099679 | 0.31855 | 4.6568 | 4.8206 | 4.7387 | 11.1359 | 0 | 0.92134
110 | 1 | 0.089727 | 0.25079 | 4.4529 | 4.6323 | 4.5426 | 10.675 | 0 | 0.92412
115 | 1 | 0.055664 | 0.42425 | 4.3527 | 4.5771 | 4.4649 | 10.4925 | 0 | 0.94051
120 | 1 | 0.32148 | 0.43145 | 4.2532 | 4.5144 | 4.3838 | 10.302 | 0 | 0.88806
125 | 1 | 0.039925 | 0.13365 | 4.1593 | 4.4631 | 4.3112 | 10.1313 | 0 | 0.90209
130 | 1 | 0.078586 | 0.49749 | 4.0692 | 4.4089 | 4.2391 | 9.9618 | 0 | 0.89687
135 | 1 | 0.30771 | 0.24001 | 3.9873 | 4.3578 | 4.1726 | 9.8055 | 0 | 0.92694
140 | 1 | -0.011696 | 0.53467 | 3.9086 | 4.3024 | 4.1055 | 9.648 | 0 | 0.91581
145 | 1 | 0.1325 | 0.23723 | 3.8575 | 4.2761 | 4.0668 | 9.557 | 0 | 1
150 | 1 | 0.0084069 | 0.48367 | 3.7695 | 4.1965 | 3.983 | 9.36 | 0 | 1
155 | 1 | 0.0025538 | 0.043637 | 3.6905 | 4.1132 | 3.9019 | 9.1694 | 0 | 1
160 | 1 | 0.19286 | 0.22239 | 3.6628 | 4.0856 | 3.8742 | 9.1044 | 0 | 0.94236
165 | 1 | 0.161 | 0.34644 | 3.6173 | 4.0089 | 3.8131 | 8.9608 | 0 | 1
170 | 1 | 0.030901 | 0.31296 | 3.5831 | 3.9541 | 3.7686 | 8.8563 | 0 | 1
175 | 1 | 0.093575 | 0.088816 | 3.5721 | 3.9209 | 3.7465 | 8.8044 | 0 | 0.89833
180 | 1 | 0.11897 | 0.13731 | 3.5433 | 3.8811 | 3.7122 | 8.7237 | 0 | 1
185 | 1 | 0.17748 | 0.25069 | 3.5378 | 3.8646 | 3.7012 | 8.6978 | 0 | 1
190 | 1 | 0.044481 | 0.087337 | 3.5364 | 3.8459 | 3.6912 | 8.6742 | 0 | 1
195 | 1 | 0.041122 | 0.24312 | 3.5332 | 3.8285 | 3.6809 | 8.65 | 0 | 1
200 | 1 | 0.18224 | 0.02637 | 3.4957 | 3.7778 | 3.6367 | 8.5463 | 0 | 0.92415
205 | 1 | 0.0934 | 0.33639 | 3.4252 | 3.6832 | 3.5542 | 8.3524 | 0 | 1
210 | 1 | 0.082577 | 0.053711 | 3.4307 | 3.6762 | 3.5534 | 8.3506 | 0 | 0.89515
215 | 1 | 0.12917 | 0.2648 | 3.4277 | 3.6612 | 3.5444 | 8.3294 | 0 | 1
220 | 1 | 0.10714 | 0.12043 | 3.3765 | 3.5693 | 3.4729 | 8.1613 | 0 | 0.96856
225 | 1 | 0.16345 | 0.096464 | 3.3272 | 3.4873 | 3.4072 | 8.0069 | 0 | 1
230 | 1 | 0.38227 | 0.16322 | 3.3244 | 3.4627 | 3.3936 | 7.9749 | 0 | 1
235 | 1 | 0.052219 | 0.20179 | 3.3088 | 3.4358 | 3.3723 | 7.9249 | 0 | 1
240 | 1 | 0.0012707 | 0.2893 | 3.3264 | 3.4382 | 3.3823 | 7.9485 | 0 | 1
245 | 1 | 0.22419 | -0.22779 | 3.3418 | 3.46 | 3.4009 | 7.9922 | 0 | 1

## G2_GA0
Gantry 2 GA0 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 2\Commissioning\Profiles\Data\20210604 0607 GA0 analysis\Fits_G0_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | theta | amplitude
70 | 1 | 0.44936 | 0.80259 | 6.1823 | 5.9364 | 6.0593 | 0 | 0.89751
75 | 1 | 0.52912 | 0.91596 | 5.8952 | 5.7329 | 5.814 | 0 | 0.91353
80 | 1 | 0.23557 | 0.47217 | 5.5999 | 5.4969 | 5.5484 | 0 | 0.89928
85 | 1 | 0.18546 | 0.31385 | 5.3961 | 5.3048 | 5.3505 | 0 | 0.95976
90 | 1 | 0.17319 | 0.53494 | 5.292 | 5.1748 | 5.2334 | 0 | 0.90179
95 | 1 | -0.051514 | 0.26072 | 5.0783 | 4.9275 | 5.0029 | 0 | 0.84643
100 | 1 | 0.01187 | 0.35938 | 5.0232 | 4.8275 | 4.9254 | 0 | 0.98899
105 | 1 | 0.25421 | 0.19398 | 4.8534 | 4.6168 | 4.7351 | 0 | 0.87102
110 | 1 | 0.19104 | 0.38503 | 4.848 | 4.5631 | 4.7055 | 0 | 0.99507
115 | 1 | 0.41119 | 0.27021 | 4.7344 | 4.415 | 4.5747 | 0 | 0.9929
120 | 1 | 0.23943 | 0.30714 | 4.6604 | 4.3285 | 4.4944 | 0 | 0.99533
125 | 1 | 0.2862 | 0.44839 | 4.5847 | 4.2389 | 4.4118 | 0 | 0.89895
130 | 1 | 0.27985 | 0.25609 | 4.4797 | 4.1375 | 4.3086 | 0 | 0.99309
135 | 1 | 0.18915 | 0.19853 | 4.3785 | 4.0544 | 4.2165 | 0 | 1
140 | 1 | 0.40676 | 0.26382 | 4.3689 | 4.0407 | 4.2048 | 0 | 0.99813
145 | 1 | -0.013728 | 0.251 | 4.2437 | 3.9345 | 4.0891 | 0 | 0.99373
150 | 1 | -0.044907 | 0.44371 | 4.2259 | 3.938 | 4.082 | 0 | 0.99888
155 | 1 | 0.11644 | 0.32012 | 4.138 | 3.8967 | 4.0173 | 0 | 0.86771
160 | 1 | 0.0081025 | 0.47788 | 4.1061 | 3.8824 | 3.9942 | 0 | 0.99927
165 | 0.99902 | 0.072831 | 0.19282 | 4.0013 | 3.8045 | 3.9029 | 0 | 0.99806
170 | 1 | 0.20038 | 0.4382 | 4.0056 | 3.8242 | 3.9149 | 0 | 0.99971
175 | 1 | 0.13961 | 0.3065 | 3.9362 | 3.7635 | 3.8499 | 0 | 0.99948
180 | 1 | 0.081769 | 0.1299 | 3.8989 | 3.7302 | 3.8145 | 0 | 0.99981
185 | 1 | 0.16703 | 0.35484 | 3.9038 | 3.7311 | 3.8175 | 0 | 1
190 | 1 | 0.28271 | 0.2011 | 3.8547 | 3.6851 | 3.7699 | 0 | 1
195 | 1 | -0.10152 | 0.21138 | 3.8544 | 3.6792 | 3.7668 | 0 | 1
200 | 1 | 0.065151 | 0.32264 | 3.8191 | 3.6451 | 3.7321 | 0 | 1
205 | 1 | 0.17798 | 0.3186 | 3.7771 | 3.6042 | 3.6906 | 0 | 1
210 | 1 | -0.16282 | 0.024926 | 3.7817 | 3.6019 | 3.6918 | 0 | 1
215 | 1 | 0.0088853 | 0.28745 | 3.7432 | 3.5666 | 3.6549 | 0 | 1
220 | 1 | 0.11351 | 0.27382 | 3.703 | 3.5465 | 3.6247 | 0 | 0.92831
225 | 1 | 0.053555 | 0.080726 | 3.6552 | 3.5269 | 3.5911 | 0 | 1
230 | 1 | 0.36212 | 0.26992 | 3.6251 | 3.5332 | 3.5791 | 0 | 1
235 | 1 | 0.17928 | 0.054737 | 3.534 | 3.4975 | 3.5158 | 0 | 1
240 | 1 | 0.20729 | 0.36965 | 3.4424 | 3.496 | 3.4692 | 0 | 1
245 | 1 | 0.14793 | 0.084976 | 3.2749 | 3.2262 | 3.2505 | 0 | 1

## G2_GA0_annual
Gantry 2 GA0 annual data
T:\Routine QA\Annual Checks\G2\2022\Profiles\Acquired Profiles\Fits_G0_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.084285 | -0.081431 | 6.1955 | 5.9148 | 6.0551 | 14.2295 | 0 | 0.97637
75 | 1 | 0.21113 | 0.037375 | 5.8891 | 5.7087 | 5.7989 | 13.6274 | 0 | 0.98713
80 | 1 | 0.19293 | -0.13809 | 5.6312 | 5.5322 | 5.5817 | 13.117 | 0 | 0.99128
85 | 1 | 0.16239 | -0.024327 | 5.3749 | 5.2976 | 5.3363 | 12.5402 | 0 | 0.98827
90 | 1 | 0.17526 | -0.11691 | 5.233 | 5.1362 | 5.1846 | 12.1838 | 0 | 0.99323
95 | 1 | 0.10348 | 0.14019 | 5.1202 | 4.9704 | 5.0453 | 11.8564 | 0 | 0.99111
100 | 1 | 0.0098823 | 0.17409 | 4.9743 | 4.7851 | 4.8797 | 11.4674 | 0 | 0.99398
105 | 1 | 0.10791 | 0.11478 | 4.8974 | 4.6681 | 4.7827 | 11.2394 | 0 | 0.99411
110 | 1 | -0.023418 | 0.019515 | 4.8289 | 4.554 | 4.6915 | 11.025 | 0 | 0.99723
115 | 0.99427 | 0.11762 | 0.14569 | 4.7156 | 4.4089 | 4.5622 | 10.7212 | 0 | 0.99403
120 | 1 | 0.1194 | 0.14124 | 4.6452 | 4.3193 | 4.4822 | 10.5332 | 0 | 0.87678
125 | 1 | -0.057531 | 0.095894 | 4.5803 | 4.2462 | 4.4132 | 10.3711 | 0 | 0.99795
130 | 1 | 0.077238 | -0.0348 | 4.508 | 4.1597 | 4.3338 | 10.1845 | 0 | 0.99802
135 | 1 | 0.034073 | 0.11546 | 4.4292 | 4.0907 | 4.2599 | 10.0109 | 0 | 0.99813
140 | 1 | -0.071713 | -0.02997 | 4.3371 | 4.013 | 4.1751 | 9.8115 | 0 | 0.99784
145 | 1 | 0.10496 | -0.10502 | 4.3003 | 3.9903 | 4.1453 | 9.7415 | 0 | 0.99854
150 | 0.99765 | 0.00087389 | 7.4707e-05 | 4.1708 | 3.8872 | 4.029 | 9.4682 | 0 | 0.99635
155 | 1 | 0.10484 | 0.068774 | 4.1579 | 3.9032 | 4.0305 | 9.4718 | 0 | 0.91827
160 | 1 | 0.13451 | 0.1558 | 4.085 | 3.8673 | 3.9761 | 9.3439 | 0 | 0.99853
165 | 1 | -0.12246 | 0.079772 | 4.0597 | 3.854 | 3.9569 | 9.2987 | 0 | 0.9989
170 | 1 | -0.15507 | -0.0037535 | 4.0052 | 3.8248 | 3.915 | 9.2003 | 0 | 0.99893
175 | 1 | -0.00057762 | 0.081394 | 3.9775 | 3.8067 | 3.8921 | 9.1464 | 0 | 0.99928
180 | 1 | -0.032645 | 0.010306 | 3.9405 | 3.7826 | 3.8616 | 9.0747 | 0 | 0.99979
185 | 1 | -0.11089 | -0.0079843 | 3.8646 | 3.7124 | 3.7885 | 8.903 | 0 | 0.99914
190 | 1 | -0.093572 | 0.10709 | 3.8692 | 3.7288 | 3.799 | 8.9276 | 0 | 1
195 | 1 | 0.091235 | 0.097039 | 3.8168 | 3.6816 | 3.7492 | 8.8106 | 0 | 1
200 | 1 | 0.034905 | 0.17774 | 3.7804 | 3.6596 | 3.72 | 8.742 | 0 | 1
205 | 1 | 0.020899 | -0.011849 | 3.7876 | 3.6533 | 3.7205 | 8.7431 | 0 | 1
210 | 1 | 0.0047581 | 0.11939 | 3.7367 | 3.6062 | 3.6715 | 8.628 | 0 | 1
215 | 1 | 0.080609 | 0.083822 | 3.6946 | 3.5694 | 3.632 | 8.5352 | 0 | 1
220 | 1 | 0.16188 | -0.14847 | 3.6541 | 3.5431 | 3.5986 | 8.4567 | 0 | 1
225 | 1 | 0.099463 | -0.031293 | 3.6086 | 3.5228 | 3.5657 | 8.3794 | 0 | 1
230 | 1 | 0.010654 | 0.018409 | 3.5828 | 3.542 | 3.5624 | 8.3717 | 0 | 1
235 | 1 | -0.00096661 | 0.036135 | 3.5057 | 3.5188 | 3.5122 | 8.2537 | 0 | 1
240 | 1 | 0.26988 | -0.023163 | 3.3866 | 3.4972 | 3.4419 | 8.0884 | 0 | 1
245 | 1 | 0.18098 | 0.15278 | 3.2711 | 3.2844 | 3.2777 | 7.7027 | 0 | 1

## G2_GA90
Gantry 2 GA90 annual QA data
T:\Routine QA\Annual Checks\G2\2022\Profiles\Acquired Profiles\Fits_G90_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 0.98606 | -0.028174 | 0.11118 | 6.0505 | 6.0959 | 6.0732 | 14.272 | 0 | 0.98025
75 | 1 | 0.045457 | 0.019527 | 5.8313 | 5.6288 | 5.7301 | 13.4656 | 0 | 0.972
80 | 1 | 0.20328 | 0.16208 | 5.6913 | 5.3862 | 5.5387 | 13.016 | 0 | 0.98854
85 | 1 | 0.062633 | -0.088921 | 5.4957 | 5.171 | 5.3333 | 12.5333 | 0 | 0.98898
90 | 1 | -0.062131 | -0.042163 | 5.3171 | 5.0382 | 5.1776 | 12.1675 | 0 | 0.98612
95 | 1 | -0.08173 | 0.02936 | 5.1291 | 4.914 | 5.0215 | 11.8006 | 0 | 0.98672
100 | 1 | 0.053826 | 0.063001 | 4.9556 | 4.8158 | 4.8857 | 11.4813 | 0 | 0.99103
105 | 1 | 0.057482 | 0.01889 | 4.7626 | 4.6978 | 4.7302 | 11.116 | 0 | 0.99389
110 | 0.99928 | 0.037102 | 0.099798 | 4.6208 | 4.6259 | 4.6233 | 10.8648 | 0 | 0.99502
115 | 1 | -0.20533 | 0.055793 | 4.4836 | 4.5376 | 4.5106 | 10.5999 | 0 | 0.99607
120 | 1 | 0.052801 | -0.059691 | 4.379 | 4.4743 | 4.4266 | 10.4026 | 0 | 0.99743
125 | 1 | 0.019455 | 0.052679 | 4.2833 | 4.3843 | 4.3338 | 10.1845 | 0 | 0.99675
130 | 0.99816 | 0.086183 | 0.071864 | 4.2227 | 4.316 | 4.2694 | 10.033 | 0 | 0.99689
135 | 1 | -0.10941 | 0.045771 | 4.1864 | 4.2652 | 4.2258 | 9.9306 | 0 | 0.99787
140 | 0.99826 | -0.072847 | -0.025027 | 4.1354 | 4.1981 | 4.1668 | 9.7919 | 0 | 0.99733
145 | 1 | -0.009323 | -0.058946 | 4.1008 | 4.1383 | 4.1196 | 9.681 | 0 | 0.99773
150 | 1 | -0.051749 | 0.03915 | 4.0693 | 4.0977 | 4.0835 | 9.5962 | 0 | 0.99816
155 | 1 | 0.033494 | 0.035553 | 4.048 | 4.0642 | 4.0561 | 9.5318 | 0 | 0.99846
160 | 1 | 0.099817 | 0.079433 | 3.976 | 3.9672 | 3.9716 | 9.3333 | 0 | 0.99817
165 | 1 | 0.021469 | -0.023313 | 3.9546 | 3.9281 | 3.9414 | 9.2622 | 0 | 0.99843
170 | 1 | 0.15794 | 0.0026418 | 3.8992 | 3.856 | 3.8776 | 9.1123 | 0 | 0.9979
175 | 1 | 0.05481 | -0.08462 | 3.8685 | 3.8114 | 3.84 | 9.0239 | 0 | 1
180 | 1 | 0.058897 | -0.016094 | 3.8555 | 3.7726 | 3.8141 | 8.9631 | 0 | 0.99711
185 | 1 | 0.082098 | -0.071526 | 3.837 | 3.7267 | 3.7819 | 8.8874 | 0 | 0.99471
190 | 1 | -0.065039 | 0.15514 | 3.843 | 3.7108 | 3.7769 | 8.8756 | 0 | 0.99497
195 | 1 | -0.113 | -0.059584 | 3.805 | 3.6571 | 3.7311 | 8.768 | 0 | 0.9854
200 | 1 | -0.10168 | 0.012907 | 3.8284 | 3.6382 | 3.7333 | 8.7733 | 0 | 0.99187
205 | 1 | -0.1158 | -0.022566 | 3.8003 | 3.5839 | 3.6921 | 8.6764 | 0 | 0.98585
210 | 1 | -0.073781 | 0.14677 | 3.7483 | 3.4945 | 3.6214 | 8.5103 | 0 | 0.95907
215 | 1 | -0.12064 | 0.032528 | 3.6866 | 3.405 | 3.5458 | 8.3326 | 0 | 0.8997
220 | 0.9833 | 0.032323 | 0.00025002 | 3.6662 | 3.3718 | 3.519 | 8.2697 | 0 | 0.92888
225 | 1 | 0.026567 | -0.048801 | 3.5465 | 3.2528 | 3.3997 | 7.9893 | 0 | 0.85597
230 | 1 | 0.067321 | 0.041884 | 3.5028 | 3.2 | 3.3514 | 7.8758 | 0 | 0.97617
235 | 0.99463 | 0.083364 | 0.11027 | 3.4096 | 3.0893 | 3.2495 | 7.6363 | 0 | 0.98286
240 | 1 | 0.15354 | -0.037615 | 3.2689 | 2.9329 | 3.1009 | 7.2871 | 0 | 0.98662
245 | 1 | -0.089174 | 0.082425 | 3.1059 | 2.723 | 2.9144 | 6.8489 | 0 | 0.99377

## G2_GA180
Gantry 2 GA180 annual QA data
T:\Routine QA\Annual Checks\G2\2022\Profiles\Acquired Profiles\Fits_G180_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | -0.015061 | -0.030281 | 6.2561 | 6.0094 | 6.1327 | 14.4119 | 0 | 0.98179
75 | 1 | -0.024352 | 0.080342 | 5.9057 | 5.7755 | 5.8406 | 13.7253 | 0 | 0.98342
80 | 1 | 0.074265 | -0.12361 | 5.5839 | 5.545 | 5.5644 | 13.0764 | 0 | 0.98263
85 | 1 | -0.12441 | 0.12286 | 5.4147 | 5.3965 | 5.4056 | 12.7031 | 0 | 0.98529
90 | 1 | -0.011392 | 0.098762 | 5.2573 | 5.2149 | 5.2361 | 12.3048 | 0 | 0.98789
95 | 1 | -0.052038 | -0.057649 | 5.0821 | 4.9976 | 5.0398 | 11.8436 | 0 | 0.98784
100 | 1 | -0.041598 | -0.06895 | 4.985 | 4.8479 | 4.9164 | 11.5536 | 0 | 0.7
105 | 1 | 0.001043 | 0.031035 | 4.869 | 4.6941 | 4.7816 | 11.2367 | 0 | 0.98987
110 | 1 | -0.21178 | -0.12063 | 4.7689 | 4.5566 | 4.6628 | 10.9575 | 0 | 0.99098
115 | 1 | -0.12377 | 0.11479 | 4.6757 | 4.4321 | 4.5539 | 10.7017 | 0 | 0.99144
120 | 1 | -0.23707 | -0.11615 | 4.5587 | 4.2865 | 4.4226 | 10.3932 | 0 | 0.99118
125 | 1 | -0.18474 | -0.066165 | 4.5344 | 4.2558 | 4.3951 | 10.3285 | 0 | 0.99426
130 | 1 | -0.17192 | -0.13003 | 4.4204 | 4.1425 | 4.2815 | 10.0615 | 0 | 0.99288
135 | 1 | 0.036357 | 0.087423 | 4.351 | 4.0822 | 4.2166 | 9.9091 | 0 | 0.994
140 | 1 | -0.25778 | -0.065417 | 4.2949 | 4.0409 | 4.1679 | 9.7945 | 0 | 0.99536
145 | 1 | -0.13023 | -0.057418 | 4.2311 | 3.9964 | 4.1137 | 9.6673 | 0 | 0.99594
150 | 1 | -0.025669 | -0.01667 | 4.1337 | 3.9309 | 4.0323 | 9.4759 | 0 | 0.99574
155 | 1 | -0.11775 | 0.071135 | 4.0696 | 3.893 | 3.9813 | 9.3561 | 0 | 0.99602
160 | 1 | -0.0017536 | 0.16279 | 4.003 | 3.8638 | 3.9334 | 9.2434 | 0 | 0.99624
165 | 1 | -0.23619 | 0.083785 | 3.9599 | 3.8491 | 3.9045 | 9.1755 | 0 | 0.9965
170 | 1 | -0.22268 | 0.12642 | 3.8801 | 3.803 | 3.8416 | 9.0276 | 0 | 0.99644
175 | 1 | -0.030732 | 0.08185 | 3.8577 | 3.7939 | 3.8258 | 8.9906 | 0 | 0.9969
180 | 1 | -0.022027 | 0.062528 | 3.8127 | 3.7677 | 3.7902 | 8.907 | 0 | 0.99705
185 | 1 | 0.03217 | -0.060003 | 3.7563 | 3.7187 | 3.7375 | 8.7832 | 0 | 0.99725
190 | 1 | -0.17593 | -0.061382 | 3.7304 | 3.699 | 3.7147 | 8.7295 | 0 | 0.99757
195 | 1 | -0.13962 | -0.15355 | 3.7033 | 3.6741 | 3.6887 | 8.6684 | 0 | 0.99777
200 | 1 | -0.18621 | -0.14578 | 3.6771 | 3.6493 | 3.6632 | 8.6086 | 0 | 0.99806
205 | 1 | -0.077752 | -0.027255 | 3.6551 | 3.6223 | 3.6387 | 8.5509 | 0 | 0.99809
210 | 1 | -0.13686 | -0.13607 | 3.6455 | 3.6095 | 3.6275 | 8.5246 | 0 | 0.99859
215 | 1 | -0.10937 | -0.085683 | 3.6118 | 3.579 | 3.5954 | 8.4492 | 0 | 0.99863
220 | 1 | 0.035957 | -0.041776 | 3.5563 | 3.5367 | 3.5465 | 8.3343 | 0 | 0.99846
225 | 1 | -0.047612 | 0.028875 | 3.5311 | 3.5431 | 3.5371 | 8.3121 | 0 | 0.94934
230 | 1 | -0.069517 | -0.04106 | 3.459 | 3.5143 | 3.4867 | 8.1937 | 0 | 0.99893
235 | 1 | 0.16095 | 0.022858 | 3.4218 | 3.5187 | 3.4703 | 8.1551 | 0 | 1
240 | 1 | 0.23057 | 0.039891 | 3.3138 | 3.4907 | 3.4023 | 7.9953 | 0 | 1
245 | 1 | 0.31073 | -0.0081307 | 3.2108 | 3.2844 | 3.2476 | 7.6318 | 0 | 0.9998

## G2_GA270
Gantry 2 GA270 annual QA data
T:\Routine QA\Annual Checks\G2\2022\Profiles\Acquired Profiles\Fits_G270_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave-fwhm | theta | amplitude
70 | 1 | 0.051529 | 0.14453 | 6.0931 | 6.1338 | 6.1135 | 14.3666 | 0 | 0.98288
75 | 1 | -0.082914 | -0.092702 | 5.898 | 5.7785 | 5.8383 | 13.7199 | 0 | 0.97952
80 | 1 | -0.27521 | 0.039329 | 5.7446 | 5.5045 | 5.6245 | 13.2176 | 0 | 0.98818
85 | 1 | -0.095747 | -0.049403 | 5.5975 | 5.3229 | 5.4602 | 12.8314 | 0 | 0.98829
90 | 1 | -0.059295 | -0.0099614 | 5.3774 | 5.1295 | 5.2535 | 12.3456 | 0 | 0.98403
95 | 1 | -0.17502 | 0.091799 | 5.2065 | 5.0248 | 5.1156 | 12.0217 | 0 | 0.99206
100 | 1 | 0.00048742 | 0.044107 | 5.0282 | 4.8985 | 4.9633 | 11.6638 | 0 | 0.99016
105 | 1 | -0.16395 | 0.0056186 | 4.8723 | 4.7993 | 4.8358 | 11.3642 | 0 | 0.7
110 | 1 | -0.21431 | 0.12516 | 4.7416 | 4.7267 | 4.7341 | 11.1252 | 0 | 0.99683
115 | 1 | -0.016327 | -0.060023 | 4.5698 | 4.6151 | 4.5924 | 10.7922 | 0 | 0.99526
120 | 0.98945 | -0.065342 | 0.055333 | 4.4289 | 4.5239 | 4.4764 | 10.5196 | 0 | 0.99271
125 | 1 | -0.21123 | 0.067537 | 4.3318 | 4.4572 | 4.3945 | 10.327 | 0 | 0.99672
130 | 1 | -0.064365 | -0.0028682 | 4.2473 | 4.391 | 4.3192 | 10.15 | 0 | 0.89753
135 | 1 | -0.074748 | 0.078718 | 4.1973 | 4.3486 | 4.2729 | 10.0414 | 0 | 0.94217
140 | 1 | -0.026888 | -0.024951 | 4.118 | 4.2775 | 4.1978 | 9.8648 | 0 | 0.99836
145 | 1 | -0.13693 | 0.14874 | 4.0295 | 4.1928 | 4.1112 | 9.6613 | 0 | 0.86642
150 | 1 | -0.22881 | 0.032537 | 3.9928 | 4.1506 | 4.0717 | 9.5685 | 0 | 0.99833
155 | 1 | -0.23476 | 0.16867 | 3.975 | 4.1142 | 4.0446 | 9.5048 | 0 | 1
160 | 1 | -0.070014 | -0.044862 | 3.9337 | 4.0606 | 3.9971 | 9.3932 | 0 | 0.99839
165 | 1 | -0.075122 | -0.046807 | 3.9023 | 4.0047 | 3.9535 | 9.2908 | 0 | 0.9983
170 | 1 | -0.01281 | 0.11746 | 3.9015 | 3.9889 | 3.9452 | 9.2712 | 0 | 0.99851
175 | 1 | -0.18119 | -0.091475 | 3.8769 | 3.9517 | 3.9143 | 9.1987 | 0 | 0.99832
180 | 1 | -0.0068198 | 0.13953 | 3.8508 | 3.9158 | 3.8833 | 9.1257 | 0 | 0.99814
185 | 1 | -0.043746 | 0.041347 | 3.7991 | 3.8839 | 3.8415 | 9.0275 | 0 | 0.99785
190 | 1 | -0.056788 | -0.062612 | 3.7128 | 3.8334 | 3.7731 | 8.8668 | 0 | 0.99663
195 | 1 | -0.1784 | 0.051151 | 3.6641 | 3.8074 | 3.7358 | 8.779 | 0 | 0.9959
200 | 1 | -0.14928 | 0.12062 | 3.6234 | 3.7729 | 3.6982 | 8.6907 | 0 | 0.99679
205 | 1 | -0.15039 | -0.12048 | 3.5552 | 3.6827 | 3.6189 | 8.5044 | 0 | 0.97273
210 | 1 | -0.15159 | 0.025415 | 3.5514 | 3.6481 | 3.5998 | 8.4595 | 0 | 0.99611
215 | 1 | -0.036383 | 0.0045378 | 3.5061 | 3.5626 | 3.5343 | 8.3057 | 0 | 0.99573
220 | 1 | -0.074636 | 0.13734 | 3.4616 | 3.4909 | 3.4762 | 8.1692 | 0 | 0.9959
225 | 1 | 0.0024519 | 0.017887 | 3.4069 | 3.4049 | 3.4059 | 8.0039 | 0 | 0.99118
230 | 1 | -0.11357 | 0.036999 | 3.3572 | 3.3028 | 3.33 | 7.8255 | 0 | 0.99285
235 | 1 | -0.059891 | 0.093918 | 3.2763 | 3.1812 | 3.2288 | 7.5876 | 0 | 0.98921
240 | 1 | -0.0014475 | -0.036065 | 3.1843 | 3.0146 | 3.0994 | 7.2837 | 0 | 0.97056
245 | 1 | 0.059952 | 0.11112 | 3.0086 | 2.8101 | 2.9094 | 6.837 | 0 | 0.9903

## G3_GA0
Gantry 3 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 3\Commissioning\Profiles\_G3 COMM PROFILES\Results_2021-12-02_4000\Fits_G0_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.40948 | 0.65358 | 6.0381 | 5.9475 | 5.9928 | 14.083 | 0 | 0.89759
75 | 1 | 0.0069836 | 0.69844 | 5.8614 | 5.8186 | 5.84 | 13.724 | 0 | 0.91779
80 | 1 | -0.18516 | 0.28879 | 5.5699 | 5.5834 | 5.5767 | 13.1051 | 0 | 0.95449
85 | 1 | 0.17223 | 0.17044 | 5.4418 | 5.4651 | 5.4535 | 12.8157 | 0 | 0.98495
90 | 1 | 0.074214 | 0.40645 | 5.2837 | 5.314 | 5.2989 | 12.4523 | 0 | 0.98835
95 | 1 | 0.13858 | 0.65536 | 5.071 | 5.0852 | 5.0781 | 11.9335 | 0 | 0.97783
100 | 1 | -0.035571 | 0.39205 | 4.9266 | 4.9116 | 4.9191 | 11.5599 | 0 | 0.96944
105 | 1 | 0.079999 | 0.2704 | 4.8618 | 4.7855 | 4.8236 | 11.3356 | 0 | 0.98955
110 | 1 | 0.18624 | 0.6529 | 4.7115 | 4.6105 | 4.661 | 10.9533 | 0 | 0.85246
115 | 1 | -0.021789 | 0.12388 | 4.6689 | 4.5129 | 4.5909 | 10.7886 | 0 | 0.99497
120 | 1 | 0.031424 | 0.26486 | 4.5622 | 4.3927 | 4.4775 | 10.522 | 0 | 0.99566
125 | 1 | -0.11845 | 0.11986 | 4.4941 | 4.3037 | 4.3989 | 10.3374 | 0 | 0.99751
130 | 1 | -0.10675 | 0.34692 | 4.3408 | 4.1574 | 4.2491 | 9.9854 | 0 | 0.8896
135 | 0.9978 | -0.069195 | 0.35844 | 4.2973 | 4.1026 | 4.2 | 9.87 | 0 | 0.87407
140 | 1 | -0.031323 | 0.093678 | 4.2679 | 4.0673 | 4.1676 | 9.7939 | 0 | 0.99903
145 | 1 | -0.22924 | 0.044735 | 4.1692 | 3.9861 | 4.0776 | 9.5825 | 0 | 0.99875
150 | 1 | -0.10519 | 0.29651 | 4.0854 | 3.9245 | 4.0049 | 9.4116 | 0 | 0.89436
155 | 1 | -0.059707 | 0.32293 | 4.0344 | 3.8748 | 3.9546 | 9.2934 | 0 | 0.99911
160 | 1 | -0.035588 | 0.32221 | 3.9636 | 3.8244 | 3.894 | 9.1509 | 0 | 0.99931
165 | 1 | -0.049526 | 0.29693 | 3.8913 | 3.7719 | 3.8316 | 9.0042 | 0 | 0.99922
170 | 1 | -0.065267 | 0.4257 | 3.8314 | 3.7293 | 3.7804 | 8.8838 | 0 | 0.99928
175 | 1 | -0.0063437 | 0.21138 | 3.7802 | 3.6837 | 3.7319 | 8.77 | 0 | 0.99936
180 | 1 | 0.083482 | 0.024714 | 3.7346 | 3.646 | 3.6903 | 8.6722 | 0 | 0.99962
185 | 1 | -0.03972 | 0.13398 | 3.6907 | 3.5945 | 3.6426 | 8.56 | 0 | 1
190 | 1 | -0.06372 | 0.0726 | 3.5828 | 3.5039 | 3.5434 | 8.3269 | 0 | 1
195 | 1 | -0.036378 | 0.018966 | 3.5681 | 3.4521 | 3.5101 | 8.2488 | 0 | 0.99909
200 | 1 | -0.040934 | 0.23899 | 3.5784 | 3.4405 | 3.5094 | 8.2472 | 0 | 0.99985
205 | 1 | 0.041657 | 0.024423 | 3.5526 | 3.3818 | 3.4672 | 8.1479 | 0 | 1
210 | 1 | 0.074268 | 0.24353 | 3.5391 | 3.3431 | 3.4411 | 8.0865 | 0 | 0.99978
215 | 1 | -0.16213 | 0.027366 | 3.5017 | 3.2888 | 3.3953 | 7.9789 | 0 | 0.99929
220 | 1 | -0.19921 | 0.12292 | 3.4759 | 3.2447 | 3.3603 | 7.8967 | 0 | 0.99926
225 | 1 | 0.19143 | 0.084366 | 3.4511 | 3.2133 | 3.3322 | 7.8307 | 0 | 0.99874
230 | 1 | 0.083205 | 0.16983 | 3.4263 | 3.2074 | 3.3168 | 7.7946 | 0 | 0.9989
235 | 1 | 0.019812 | 0.16929 | 3.3717 | 3.1758 | 3.2738 | 7.6933 | 0 | 0.99893
240 | 1 | 0.067405 | -0.044818 | 3.27 | 3.1459 | 3.208 | 7.5387 | 0 | 0.90622
245 | 1 | -0.14398 | 0.3933 | 3.0775 | 2.9869 | 3.0322 | 7.1256 | 0 | 0.99913

## G3_GA90
Gantry 3 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 3\Commissioning\Profiles\_G3 COMM PROFILES\Results_2021-12-15_4000\Fits_G90_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.40544 | 0.66507 | 5.9227 | 6.005 | 5.9639 | 14.0151 | 0 | 0.90209
75 | 1 | 0.17585 | 0.36733 | 5.8201 | 5.7969 | 5.8085 | 13.65 | 0 | 0.90075
80 | 1 | 0.41616 | 0.65454 | 5.678 | 5.556 | 5.617 | 13.2 | 0 | 0.97583
85 | 1 | -0.0093972 | 0.23142 | 5.5112 | 5.3481 | 5.4297 | 12.7597 | 0 | 0.89766
90 | 1 | 0.050672 | 0.47774 | 5.3514 | 5.1737 | 5.2625 | 12.367 | 0 | 0.9749
95 | 1 | 0.25877 | 0.25655 | 5.1678 | 5.0063 | 5.0871 | 11.9546 | 0 | 0.9041
100 | 1 | 0.04784 | 0.011352 | 5.0581 | 4.9223 | 4.9902 | 11.727 | 0 | 0.98801
105 | 1 | 0.18695 | 0.28577 | 4.9055 | 4.8092 | 4.8574 | 11.4149 | 0 | 0.83434
110 | 1 | 0.10844 | 0.14374 | 4.6373 | 4.5923 | 4.6148 | 10.8448 | 0 | 0.7
115 | 1 | 0.31741 | 0.34192 | 4.5911 | 4.5609 | 4.576 | 10.7536 | 0 | 0.99231
120 | 1 | 0.059571 | 0.41525 | 4.5121 | 4.5146 | 4.5134 | 10.6064 | 0 | 0.91559
125 | 1 | 0.20868 | 0.34401 | 4.3934 | 4.4059 | 4.3997 | 10.3392 | 0 | 0.99711
130 | 1 | 0.31988 | 0.12255 | 4.3319 | 4.3447 | 4.3383 | 10.195 | 0 | 0.99879
135 | 1 | 0.024626 | -0.17526 | 4.2075 | 4.2188 | 4.2132 | 9.9009 | 0 | 0.90549
140 | 1 | 0.26984 | 0.34175 | 4.1834 | 4.1828 | 4.1831 | 9.8303 | 0 | 0.99913
145 | 1 | 0.20605 | 0.078233 | 4.1169 | 4.0926 | 4.1048 | 9.6463 | 0 | 0.99898
150 | 1 | -0.00608 | 0.65868 | 4.0424 | 4.0058 | 4.0241 | 9.4566 | 0 | 1
155 | 1 | 0.38688 | 0.22025 | 4.0223 | 3.9484 | 3.9853 | 9.3655 | 0 | 0.99955
160 | 1 | 0.17832 | 0.17695 | 3.9721 | 3.8569 | 3.9145 | 9.1991 | 0 | 0.99944
165 | 1 | 0.15038 | 0.26888 | 3.8979 | 3.7669 | 3.8324 | 9.0062 | 0 | 1
170 | 1 | 0.36257 | 0.27643 | 3.9286 | 3.7745 | 3.8515 | 9.0511 | 0 | 1
175 | 1 | -0.15044 | 0.26856 | 3.8995 | 3.7296 | 3.8146 | 8.9642 | 0 | 1
180 | 1 | 0.065174 | -0.020284 | 3.8313 | 3.637 | 3.7342 | 8.7754 | 0 | 1
185 | 1 | 0.24157 | 0.30095 | 3.8237 | 3.6167 | 3.7202 | 8.7425 | 0 | 1
190 | 1 | -0.12918 | 0.046583 | 3.7777 | 3.5604 | 3.669 | 8.6222 | 0 | 1
195 | 1 | -0.023912 | 0.025505 | 3.7461 | 3.5266 | 3.6364 | 8.5454 | 0 | 1
200 | 1 | 0.23099 | 0.28806 | 3.6778 | 3.4796 | 3.5787 | 8.41 | 0 | 1
205 | 1 | 0.010459 | 0.11103 | 3.6708 | 3.5033 | 3.5871 | 8.4296 | 0 | 1
210 | 1 | -0.0061329 | -0.021674 | 3.6027 | 3.4778 | 3.5403 | 8.3197 | 0 | 1
215 | 1 | -0.11935 | -0.051654 | 3.5488 | 3.4617 | 3.5053 | 8.2374 | 0 | 1
220 | 1 | -0.061682 | 0.16538 | 3.5235 | 3.4893 | 3.5064 | 8.2402 | 0 | 1
225 | 1 | -0.064001 | 0.030494 | 3.4615 | 3.4779 | 3.4697 | 8.1538 | 0 | 0.89144
230 | 1 | 0.26141 | 0.16115 | 3.4149 | 3.485 | 3.4499 | 8.1073 | 0 | 1
235 | 1 | 0.26107 | 0.1349 | 3.3196 | 3.4346 | 3.3771 | 7.9362 | 0 | 1
240 | 1 | 0.21102 | 0.37367 | 3.2657 | 3.3823 | 3.324 | 7.8114 | 0 | 1
245 | 1 | 0.0090575 | 0.064853 | 3.0997 | 3.2371 | 3.1684 | 7.4457 | 0 | 1

## G3_GA180
Gantry 3 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 3\Commissioning\Profiles\_G3 COMM PROFILES\Results_2021-12-15_3000\Fits_G180_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.35564 | 0.63537 | 5.931 | 6.0243 | 5.9776 | 14.0475 | 0 | 0.9069
75 | 1 | 0.092168 | 0.33146 | 5.7041 | 5.7998 | 5.752 | 13.5171 | 0 | 0.96547
80 | 1 | -0.16022 | 0.19316 | 5.5436 | 5.5934 | 5.5685 | 13.086 | 0 | 0.97745
85 | 1 | 0.35206 | 0.65676 | 5.3945 | 5.4328 | 5.4136 | 12.7221 | 0 | 0.91908
90 | 1 | 0.40574 | 0.38591 | 5.2012 | 5.2506 | 5.2259 | 12.2809 | 0 | 0.97951
95 | 1 | -0.10117 | 0.55334 | 5.0448 | 5.1065 | 5.0757 | 11.9278 | 0 | 0.87209
100 | 1 | 0.076595 | 0.25371 | 4.9413 | 5.0253 | 4.9833 | 11.7108 | 0 | 0.99288
105 | 1 | 0.10077 | 0.6947 | 4.7486 | 4.8671 | 4.8078 | 11.2984 | 0 | 0.87413
110 | 1 | -0.11135 | 0.0033153 | 4.6344 | 4.7907 | 4.7126 | 11.0746 | 0 | 0.99411
115 | 1 | 0.12837 | 0.30239 | 4.4969 | 4.6746 | 4.5857 | 10.7764 | 0 | 0.99434
120 | 1 | 0.099169 | 0.30403 | 4.3876 | 4.5914 | 4.4895 | 10.5504 | 0 | 0.99608
125 | 1 | 0.03215 | 0.48446 | 4.2711 | 4.4793 | 4.3752 | 10.2818 | 0 | 0.99476
130 | 1 | 0.13681 | 0.17337 | 4.1738 | 4.4005 | 4.2872 | 10.0749 | 0 | 0.99597
135 | 1 | 0.39069 | 0.43593 | 4.0702 | 4.2891 | 4.1797 | 9.8222 | 0 | 1
140 | 1 | 0.074439 | 0.21664 | 4.0104 | 4.2316 | 4.121 | 9.6844 | 0 | 0.99733
145 | 1 | 0.072342 | 0.21325 | 3.9184 | 4.1198 | 4.0191 | 9.445 | 0 | 0.91739
150 | 1 | 0.018446 | 0.21912 | 3.8877 | 4.0783 | 3.983 | 9.3601 | 0 | 0.99744
155 | 1 | 0.074733 | 0.38464 | 3.8397 | 4.0127 | 3.9262 | 9.2267 | 0 | 0.93078
160 | 1 | 0.06942 | 0.34943 | 3.7629 | 3.9062 | 3.8345 | 9.0112 | 0 | 0.99494
165 | 1 | -0.015845 | 0.20717 | 3.7493 | 3.87 | 3.8096 | 8.9526 | 0 | 0.94243
170 | 1 | -0.014426 | 0.42762 | 3.695 | 3.7944 | 3.7447 | 8.8001 | 0 | 0.99804
175 | 1 | 0.088708 | 0.37042 | 3.6647 | 3.7597 | 3.7122 | 8.7237 | 0 | 0.99824
180 | 1 | -0.088201 | 0.29218 | 3.6103 | 3.7007 | 3.6555 | 8.5904 | 0 | 0.99765
185 | 1 | 0.0079358 | 0.19043 | 3.5401 | 3.6472 | 3.5936 | 8.445 | 0 | 0.99725
190 | 1 | 0.028619 | 0.21635 | 3.5034 | 3.6203 | 3.5619 | 8.3704 | 0 | 0.99856
195 | 1 | 0.23055 | 0.16679 | 3.4736 | 3.611 | 3.5423 | 8.3244 | 0 | 0.99914
200 | 1 | 0.12241 | -0.0055145 | 3.4051 | 3.5631 | 3.4841 | 8.1876 | 0 | 0.99659
205 | 1 | 0.040101 | 0.12734 | 3.3384 | 3.5184 | 3.4284 | 8.0567 | 0 | 0.98861
210 | 1 | 0.13248 | 0.086798 | 3.3203 | 3.5331 | 3.4267 | 8.0527 | 0 | 0.99859
215 | 1 | -0.019579 | 0.36658 | 3.2799 | 3.5074 | 3.3936 | 7.975 | 0 | 0.99573
220 | 1 | 0.025467 | 0.1211 | 3.2209 | 3.4622 | 3.3415 | 7.8526 | 0 | 0.97886
225 | 1 | 0.17968 | 0.16659 | 3.2022 | 3.4537 | 3.3279 | 7.8207 | 0 | 1
230 | 1 | 0.16181 | 0.17134 | 3.1709 | 3.4205 | 3.2957 | 7.7449 | 0 | 1
235 | 1 | 0.40281 | 0.21147 | 3.1506 | 3.373 | 3.2618 | 7.6652 | 0 | 0.99516
240 | 1 | 0.16121 | 0.081461 | 3.114 | 3.2561 | 3.1851 | 7.4849 | 0 | 0.99468
245 | 1 | 0.071289 | 0.24502 | 2.9594 | 3.0596 | 3.0095 | 7.0723 | 0 | 0.9959

## G3_GA270
Gantry 3 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 3\Commissioning\Profiles\_G3 COMM PROFILES\Results_2021-12-15_4000\Fits_G270_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.080951 | 0.56706 | 5.9136 | 6.0004 | 5.957 | 13.9989 | 0 | 0.90212
75 | 1 | 0.15052 | 0.58342 | 5.8089 | 5.7962 | 5.8025 | 13.636 | 0 | 0.91004
80 | 1 | 0.13039 | 0.051704 | 5.6057 | 5.5157 | 5.5607 | 13.0676 | 0 | 1
85 | 1 | -0.16436 | 0.79311 | 5.4948 | 5.3227 | 5.4088 | 12.7106 | 0 | 0.9144
90 | 1 | 0.11747 | 0.61252 | 5.39 | 5.2037 | 5.2968 | 12.4475 | 0 | 0.89574
95 | 1 | 0.081393 | 0.36815 | 5.1983 | 5.0442 | 5.1213 | 12.035 | 0 | 0.88583
100 | 1 | 0.13431 | 0.36636 | 5.0704 | 4.9406 | 5.0055 | 11.7629 | 0 | 0.8841
105 | 1 | 0.24252 | 0.66414 | 4.8986 | 4.8058 | 4.8522 | 11.4026 | 0 | 0.83016
110 | 1 | -0.10451 | 0.36083 | 4.7944 | 4.7207 | 4.7576 | 11.1803 | 0 | 0.84812
115 | 1 | -0.2493 | 0.1347 | 4.6266 | 4.5958 | 4.6112 | 10.8364 | 0 | 0.88175
120 | 1 | -0.31407 | 0.12907 | 4.4932 | 4.4979 | 4.4956 | 10.5646 | 0 | 0.99731
125 | 1 | -0.26802 | 0.41095 | 4.371 | 4.3977 | 4.3844 | 10.3033 | 0 | 0.88225
130 | 1 | 0.10705 | 0.39552 | 4.2813 | 4.3325 | 4.3069 | 10.1212 | 0 | 0.93648
135 | 1 | -0.094796 | 0.38622 | 4.2057 | 4.2683 | 4.237 | 9.957 | 0 | 0.99928
140 | 1 | 0.021516 | 0.38709 | 4.1873 | 4.2503 | 4.2188 | 9.9142 | 0 | 1
145 | 1 | -0.052739 | 0.32966 | 4.0974 | 4.141 | 4.1192 | 9.68 | 0 | 1
150 | 1 | 0.0087209 | 0.14236 | 4.0295 | 4.0515 | 4.0405 | 9.4952 | 0 | 0.94628
155 | 1 | -0.12079 | 0.36567 | 3.9974 | 3.9723 | 3.9849 | 9.3644 | 0 | 1
160 | 1 | -0.11798 | 0.39667 | 3.9648 | 3.8893 | 3.927 | 9.2286 | 0 | 0.93696
165 | 1 | -0.16014 | 0.36688 | 3.95 | 3.8376 | 3.8938 | 9.1504 | 0 | 1
170 | 1 | -0.099235 | 0.11328 | 3.9235 | 3.7946 | 3.8591 | 9.0688 | 0 | 1
175 | 1 | -0.096824 | 0.12375 | 3.8659 | 3.7178 | 3.7919 | 8.9109 | 0 | 1
180 | 1 | 0.18785 | 0.1222 | 3.8921 | 3.7156 | 3.8039 | 8.9391 | 0 | 1
185 | 1 | 0.048684 | 0.38209 | 3.843 | 3.6559 | 3.7495 | 8.8113 | 0 | 1
190 | 1 | 0.008261 | 0.06498 | 3.8123 | 3.6025 | 3.7074 | 8.7124 | 0 | 1
195 | 1 | 0.018796 | 0.17196 | 3.8231 | 3.6006 | 3.7119 | 8.7229 | 0 | 1
200 | 1 | 0.072441 | 0.34736 | 3.801 | 3.5835 | 3.6922 | 8.6767 | 0 | 1
205 | 1 | 0.24925 | 0.21764 | 3.7334 | 3.56 | 3.6467 | 8.5698 | 0 | 1
210 | 1 | 0.19147 | 0.36195 | 3.677 | 3.5677 | 3.6224 | 8.5126 | 0 | 1
215 | 1 | 0.16005 | -0.021822 | 3.5779 | 3.541 | 3.5594 | 8.3647 | 0 | 1
220 | 1 | 0.1391 | 0.055486 | 3.4648 | 3.5175 | 3.4911 | 8.2042 | 0 | 1
225 | 1 | 0.179 | 0.30536 | 3.4004 | 3.5131 | 3.4568 | 8.1234 | 0 | 1
230 | 1 | 0.41012 | 0.30592 | 3.3236 | 3.4972 | 3.4104 | 8.0144 | 0 | 1
235 | 1 | 0.31018 | 0.26372 | 3.3257 | 3.5142 | 3.4199 | 8.0369 | 0 | 1
240 | 1 | 0.14753 | 0.17786 | 3.2823 | 3.4301 | 3.3562 | 7.8871 | 0 | 1
245 | 1 | 0.028192 | 0.24438 | 3.0305 | 3.2727 | 3.1516 | 7.4063 | 0 | 1

## G4_GA0
Gantry 4 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 4\Commissioning\Spot Profiles\_GaussianFittingResults\Fits_G0_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 0.99618 | 0.35919 | 0.17997 | 5.7765 | 5.9533 | 5.8649 | 13.7825 | 0 | 0.94874
75 | 1 | -0.11045 | 0.42802 | 5.526 | 5.6588 | 5.5924 | 13.1422 | 0 | 0.93507
80 | 1 | 0.13728 | 0.53976 | 5.3518 | 5.5099 | 5.4309 | 12.7625 | 0 | 0.9819
85 | 1 | -0.035076 | -0.28177 | 5.1666 | 5.3442 | 5.2554 | 12.3502 | 0 | 0.98349
90 | 1 | 0.15786 | 0.21549 | 5.0016 | 5.217 | 5.1093 | 12.0068 | 0 | 0.98957
95 | 1 | -0.15876 | 0.26352 | 4.834 | 5.0843 | 4.9592 | 11.654 | 0 | 0.99168
100 | 1 | 0.020258 | 0.0014376 | 4.6782 | 4.9467 | 4.8124 | 11.3092 | 0 | 0.99281
105 | 1 | 0.22528 | 0.30813 | 4.5526 | 4.8193 | 4.686 | 11.012 | 0 | 0.99535
110 | 1 | -0.012086 | 0.32333 | 4.4219 | 4.6841 | 4.553 | 10.6995 | 0 | 0.99575
115 | 1 | 0.17932 | -0.00077478 | 4.335 | 4.5681 | 4.4515 | 10.4611 | 0 | 0.99738
120 | 1 | 0.0083768 | 0.24009 | 4.2576 | 4.494 | 4.3758 | 10.2831 | 0 | 0.99888
125 | 1 | 0.070215 | 0.25165 | 4.1106 | 4.3519 | 4.2313 | 9.9435 | 0 | 1
130 | 1 | 0.21704 | 0.1892 | 4.0924 | 4.2941 | 4.1933 | 9.8542 | 0 | 0.89238
135 | 1 | -0.12082 | 0.25158 | 4.042 | 4.2358 | 4.1389 | 9.7264 | 0 | 0.88416
140 | 1 | -0.12356 | 0.13773 | 4.0018 | 4.1817 | 4.0917 | 9.6156 | 0 | 0.99937
145 | 1 | -0.059815 | 0.45851 | 3.9083 | 4.1077 | 4.008 | 9.4188 | 0 | 1
150 | 1 | -0.13318 | 0.40263 | 3.8867 | 4.067 | 3.9768 | 9.3456 | 0 | 0.913
155 | 1 | -0.25742 | 0.12778 | 3.8388 | 4.0211 | 3.9299 | 9.2354 | 0 | 1
160 | 1 | -0.049657 | 0.38691 | 3.7686 | 3.9699 | 3.8692 | 9.0927 | 0 | 1
165 | 1 | -0.1989 | 0.12782 | 3.7059 | 3.9142 | 3.8101 | 8.9537 | 0 | 1
170 | 1 | -0.0025622 | 0.13668 | 3.6832 | 3.9043 | 3.7937 | 8.9153 | 0 | 1
175 | 1 | 0.0089413 | -0.05013 | 3.6164 | 3.8308 | 3.7236 | 8.7505 | 0 | 1
180 | 1 | -0.097066 | 0.25527 | 3.5695 | 3.7864 | 3.678 | 8.6433 | 0 | 1
185 | 1 | -0.0925 | -0.13508 | 3.5369 | 3.7279 | 3.6324 | 8.5361 | 0 | 1
190 | 1 | 0.068986 | 0.16275 | 3.5063 | 3.6864 | 3.5964 | 8.4515 | 0 | 1
195 | 1 | 0.089609 | 0.14727 | 3.4888 | 3.6516 | 3.5702 | 8.39 | 0 | 1
200 | 1 | -0.1228 | 0.13779 | 3.4506 | 3.5919 | 3.5212 | 8.2749 | 0 | 1
205 | 1 | -0.065797 | 0.36662 | 3.3778 | 3.4988 | 3.4383 | 8.0799 | 0 | 1
210 | 1 | -0.19633 | 0.17905 | 3.3589 | 3.4542 | 3.4066 | 8.0054 | 0 | 1
215 | 1 | 0.09548 | 0.23592 | 3.3204 | 3.3944 | 3.3574 | 7.8899 | 0 | 1
220 | 1 | -0.17272 | 0.34537 | 3.2978 | 3.331 | 3.3144 | 7.7889 | 0 | 1
225 | 1 | -0.15423 | 0.3908 | 3.2773 | 3.2687 | 3.273 | 7.6916 | 0 | 1
230 | 1 | -0.11706 | 0.50675 | 3.2892 | 3.23 | 3.2596 | 7.6601 | 0 | 0.99996
235 | 1 | -0.095793 | 0.09584 | 3.2934 | 3.1978 | 3.2456 | 7.6271 | 0 | 0.99948
240 | 1 | -0.043505 | 0.37463 | 3.2258 | 3.1706 | 3.1982 | 7.5158 | 0 | 0.99916
245 | 1 | 0.080355 | 0.1357 | 3.032 | 2.9691 | 3.0005 | 7.0513 | 0 | 0.9988

## G4_GA90
Gantry 4 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 4\Commissioning\Spot Profiles\_GaussianFittingResults\Fits_G90_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.088315 | 0.67569 | 5.8449 | 5.9234 | 5.8841 | 13.8277 | 0 | 0.7
75 | 1 | -0.059583 | 0.66483 | 5.6037 | 5.7371 | 5.6704 | 13.3254 | 0 | 0.90262
80 | 1 | 0.42635 | 0.38898 | 5.4361 | 5.5378 | 5.487 | 12.8944 | 0 | 0.96029
85 | 1 | 0.16757 | 0.37931 | 5.3068 | 5.3696 | 5.3382 | 12.5448 | 0 | 0.98894
90 | 1 | 0.067476 | 0.1254 | 5.0799 | 5.0962 | 5.088 | 11.9569 | 0 | 0.88102
95 | 1 | 0.26614 | 0.10614 | 5.0067 | 4.9912 | 4.999 | 11.7476 | 0 | 0.7
100 | 1 | 0.2633 | 0.36104 | 4.8665 | 4.8323 | 4.8494 | 11.396 | 0 | 0.9912
105 | 1 | 0.45903 | 0.39023 | 4.7237 | 4.6765 | 4.7001 | 11.0453 | 0 | 0.90549
110 | 1 | 0.12489 | 0.40193 | 4.5847 | 4.5407 | 4.5627 | 10.7224 | 0 | 0.89119
115 | 1 | 0.29127 | 0.33159 | 4.4992 | 4.4626 | 4.4809 | 10.5301 | 0 | 0.92968
120 | 1 | 0.17486 | 0.35704 | 4.3573 | 4.3246 | 4.341 | 10.2013 | 0 | 1
125 | 1 | 0.093941 | 0.31462 | 4.3309 | 4.3064 | 4.3187 | 10.1488 | 0 | 0.99881
130 | 1 | 0.039718 | 0.29263 | 4.2399 | 4.2099 | 4.2249 | 9.9285 | 0 | 0.99867
135 | 1 | -0.022994 | 0.27141 | 4.1602 | 4.13 | 4.1451 | 9.741 | 0 | 0.91002
140 | 1 | -0.12442 | -0.014751 | 4.1227 | 4.0922 | 4.1074 | 9.6524 | 0 | 0.99935
145 | 1 | -0.020404 | 0.26212 | 4.0614 | 4.0282 | 4.0448 | 9.5052 | 0 | 0.99998
150 | 1 | 0.12172 | 0.21245 | 3.9964 | 3.9569 | 3.9767 | 9.3452 | 0 | 1
155 | 1 | 0.1807 | 0.21895 | 3.9828 | 3.9042 | 3.9435 | 9.2672 | 0 | 1
160 | 1 | -0.15222 | -0.061513 | 3.9577 | 3.8548 | 3.9063 | 9.1797 | 0 | 1
165 | 1 | -0.046468 | 0.21305 | 3.9233 | 3.8026 | 3.863 | 9.078 | 0 | 1
170 | 1 | -0.046513 | 0.38594 | 3.8517 | 3.725 | 3.7883 | 8.9026 | 0 | 1
175 | 1 | 0.14571 | 0.20787 | 3.8061 | 3.6663 | 3.7362 | 8.7801 | 0 | 1
180 | 1 | 0.10743 | 0.17941 | 3.7681 | 3.6242 | 3.6961 | 8.6859 | 0 | 1
185 | 1 | -0.016345 | 0.18818 | 3.7113 | 3.5829 | 3.6471 | 8.5707 | 0 | 1
190 | 1 | 0.1063 | 0.47504 | 3.629 | 3.4989 | 3.5639 | 8.3752 | 0 | 1
195 | 1 | -0.098445 | 0.14032 | 3.6074 | 3.4892 | 3.5483 | 8.3385 | 0 | 1
200 | 1 | 0.11127 | 0.42763 | 3.5346 | 3.4302 | 3.4824 | 8.1836 | 0 | 1
205 | 1 | -0.073973 | 0.033621 | 3.5201 | 3.4431 | 3.4816 | 8.1818 | 0 | 1
210 | 1 | 0.19276 | 0.41157 | 3.4383 | 3.396 | 3.4171 | 8.0303 | 0 | 1
215 | 1 | 0.30896 | 0.35053 | 3.3876 | 3.3961 | 3.3918 | 7.9708 | 0 | 1
220 | 1 | 0.35865 | 0.11104 | 3.3407 | 3.4102 | 3.3755 | 7.9324 | 0 | 1
225 | 1 | 0.14276 | -0.11934 | 3.2823 | 3.4102 | 3.3462 | 7.8636 | 0 | 1
230 | 1 | 0.42376 | 0.13776 | 3.2338 | 3.4208 | 3.3273 | 7.8191 | 0 | 1
235 | 1 | 0.29483 | 0.32001 | 3.2062 | 3.4193 | 3.3128 | 7.785 | 0 | 0.99998
240 | 1 | 0.20165 | 0.2228 | 3.199 | 3.3702 | 3.2846 | 7.7188 | 0 | 1
245 | 1 | 0.058688 | 0.18756 | 2.9905 | 3.1048 | 3.0477 | 7.162 | 0 | 0.99866

## G4_GA180
Gantry 4 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 4\Commissioning\Spot Profiles\_GaussianFittingResults\Fits_G180_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | -0.098185 | 0.28119 | 5.8015 | 5.9753 | 5.8884 | 13.8377 | 0 | 0.97556
75 | 1 | -0.012562 | 0.086812 | 5.5839 | 5.6968 | 5.6403 | 13.2548 | 0 | 0.98057
80 | 1 | 0.35237 | 0.29452 | 5.4448 | 5.5891 | 5.517 | 12.9648 | 0 | 0.99375
85 | 1 | 0.44799 | 0.60774 | 5.2357 | 5.4186 | 5.3271 | 12.5188 | 0 | 0.9061
90 | 1 | 0.13011 | 0.087309 | 5.0141 | 5.2369 | 5.1255 | 12.045 | 0 | 0.9929
95 | 1 | 0.048854 | -0.0041032 | 4.8485 | 5.0953 | 4.9719 | 11.684 | 0 | 0.99411
100 | 1 | -0.09423 | 0.50845 | 4.6878 | 4.9615 | 4.8247 | 11.338 | 0 | 0.99416
105 | 1 | 0.32668 | 0.48101 | 4.4871 | 4.7653 | 4.6262 | 10.8715 | 0 | 0.8795
110 | 1 | -0.038387 | 0.4795 | 4.4121 | 4.6898 | 4.5509 | 10.6947 | 0 | 0.99549
115 | 1 | 0.17966 | 0.20752 | 4.3184 | 4.5896 | 4.454 | 10.4669 | 0 | 0.94743
120 | 1 | -0.056074 | 0.31472 | 4.2024 | 4.4635 | 4.3329 | 10.1824 | 0 | 0.91828
125 | 1 | 0.19587 | 0.22138 | 4.164 | 4.4167 | 4.2904 | 10.0824 | 0 | 0.90874
130 | 1 | 0.04581 | 0.20195 | 4.0937 | 4.3339 | 4.2138 | 9.9025 | 0 | 0.9983
135 | 1 | -0.024934 | -0.11921 | 4.025 | 4.2583 | 4.1417 | 9.7329 | 0 | 0.99859
140 | 1 | 0.092389 | 0.20531 | 3.9668 | 4.1917 | 4.0792 | 9.5862 | 0 | 0.90961
145 | 1 | 0.057005 | 0.16577 | 3.9793 | 4.146 | 4.0626 | 9.5472 | 0 | 1
150 | 1 | -0.10394 | 0.1782 | 3.8631 | 4.0859 | 3.9745 | 9.3401 | 0 | 0.91036
155 | 1 | -0.1487 | 0.24439 | 3.8179 | 4.048 | 3.933 | 9.2425 | 0 | 1
160 | 1 | 0.19655 | 0.18088 | 3.7763 | 4.0094 | 3.8928 | 9.1482 | 0 | 1
165 | 1 | 0.064029 | 0.3053 | 3.7229 | 3.9612 | 3.8421 | 9.0289 | 0 | 1
170 | 1 | 0.10688 | 0.22423 | 3.636 | 3.8717 | 3.7538 | 8.8215 | 0 | 1
175 | 1 | -0.094043 | 0.046604 | 3.6132 | 3.8463 | 3.7297 | 8.7649 | 0 | 1
180 | 1 | -0.0024752 | -0.023426 | 3.5719 | 3.8025 | 3.6872 | 8.6649 | 0 | 1
185 | 1 | -0.13056 | 0.20138 | 3.5396 | 3.7478 | 3.6437 | 8.5627 | 0 | 1
190 | 1 | -0.16423 | 0.041296 | 3.5231 | 3.7034 | 3.6133 | 8.4912 | 0 | 1
195 | 1 | 0.076144 | 0.28989 | 3.4757 | 3.6373 | 3.5565 | 8.3578 | 0 | 1
200 | 1 | 0.11739 | 0.31828 | 3.474 | 3.6114 | 3.5427 | 8.3253 | 0 | 1
205 | 1 | -0.13018 | 0.035209 | 3.4244 | 3.5339 | 3.4791 | 8.1759 | 0 | 1
210 | 1 | 0.17587 | 0.053642 | 3.3812 | 3.4756 | 3.4284 | 8.0567 | 0 | 1
215 | 1 | -0.074168 | -0.024065 | 3.3356 | 3.4047 | 3.3702 | 7.9199 | 0 | 1
220 | 1 | -0.056282 | 0.24434 | 3.3362 | 3.3709 | 3.3535 | 7.8808 | 0 | 1
225 | 1 | -0.036706 | -0.057201 | 3.2877 | 3.2723 | 3.28 | 7.708 | 0 | 1
230 | 1 | -0.010348 | -0.22059 | 3.306 | 3.2454 | 3.2757 | 7.6979 | 0 | 1
235 | 1 | -0.099802 | 0.18109 | 3.3031 | 3.1986 | 3.2508 | 7.6395 | 0 | 1
240 | 1 | 0.14304 | 0.27682 | 3.2249 | 3.1455 | 3.1852 | 7.4852 | 0 | 1
245 | 1 | 0.054329 | 0.2249 | 3.0822 | 2.9752 | 3.0287 | 7.1175 | 0 | 0.99969

## G4_GA270
Gantry 4 commissioning data
C:\Users\KAWCHUNG\NHS\(Canc) Radiotherapy - PBT Physics Team - PBT Physics Team\QAandCommissioning\Gantry 4\Commissioning\Spot Profiles\_GaussianFittingResults\Fits_G270_RS0_Dist0.xlsx
Single Gaussian Fit
Energy | amplitude | x_mean | y_mean | x_stddev | y_stddev | ave_stddev | ave_fwhm | theta | amplitude
70 | 1 | 0.30791 | 0.13766 | 5.8059 | 5.7898 | 5.7979 | 13.625 | 0 | 0.96578
75 | 1 | 0.062793 | 0.63008 | 5.5764 | 5.5813 | 5.5789 | 13.1104 | 0 | 0.98645
80 | 1 | 0.24469 | 0.22126 | 5.3753 | 5.3477 | 5.3615 | 12.5996 | 0 | 0.97834
85 | 1 | 0.1652 | 0.34796 | 5.243 | 5.1868 | 5.2149 | 12.255 | 0 | 0.88677
90 | 1 | 0.16071 | 0.19255 | 5.0722 | 4.9857 | 5.0289 | 11.818 | 0 | 0.85022
95 | 1 | 0.36143 | 0.17276 | 4.9999 | 4.8782 | 4.9391 | 11.6068 | 0 | 0.84657
100 | 1 | 0.18925 | 0.1 | 4.8821 | 4.7567 | 4.8194 | 11.3255 | 0 | 0.90045
105 | 0.99704 | -0.12864 | -0.066751 | 4.6467 | 4.5431 | 4.5949 | 10.7981 | 0 | 0.92301
110 | 1 | -0.23162 | -0.11798 | 4.5908 | 4.4671 | 4.529 | 10.6431 | 0 | 0.90999
115 | 1 | 0.071695 | -0.13425 | 4.5094 | 4.3845 | 4.4469 | 10.4503 | 0 | 0.9986
120 | 1 | 0.15119 | -0.037954 | 4.4043 | 4.2805 | 4.3424 | 10.2047 | 0 | 0.88788
125 | 1 | -0.1107 | 0.13611 | 4.3784 | 4.2515 | 4.3149 | 10.1401 | 0 | 0.9332
130 | 1 | -0.13033 | 0.11666 | 4.2936 | 4.1632 | 4.2284 | 9.9367 | 0 | 0.95475
135 | 1 | 0.18584 | 0.11774 | 4.3009 | 4.1566 | 4.2288 | 9.9376 | 0 | 1
140 | 1 | -0.066283 | 0.10822 | 4.2073 | 4.0595 | 4.1334 | 9.7136 | 0 | 1
145 | 1 | -0.14196 | 0.37113 | 4.1553 | 4.026 | 4.0907 | 9.613 | 0 | 1
150 | 1 | -0.15399 | 0.094225 | 4.0914 | 3.9752 | 4.0333 | 9.4782 | 0 | 1
155 | 1 | 0.092397 | 0.073851 | 4.043 | 3.9175 | 3.9802 | 9.3535 | 0 | 1
160 | 1 | 0.17799 | 0.069221 | 4.0132 | 3.9127 | 3.963 | 9.3129 | 0 | 0.88334
165 | 1 | -0.06159 | 0.33861 | 3.9718 | 3.9001 | 3.9359 | 9.2495 | 0 | 1
170 | 1 | -0.20195 | -0.22998 | 3.9095 | 3.8546 | 3.8821 | 9.1229 | 0 | 1
175 | 1 | -0.096948 | 0.32396 | 3.8792 | 3.8269 | 3.853 | 9.0547 | 0 | 1
180 | 1 | -0.26163 | 0.042711 | 3.8662 | 3.8004 | 3.8333 | 9.0082 | 0 | 1
185 | 1 | -0.28597 | 0.026128 | 3.7932 | 3.7292 | 3.7612 | 8.8388 | 0 | 1
190 | 1 | 0.035192 | 0.03152 | 3.8116 | 3.712 | 3.7618 | 8.8402 | 0 | 1
195 | 1 | 0.038809 | 0.29756 | 3.7603 | 3.6321 | 3.6962 | 8.6861 | 0 | 1
200 | 1 | -0.27069 | 0.29373 | 3.7796 | 3.5812 | 3.6804 | 8.649 | 0 | 1
205 | 1 | -0.01586 | 0.28497 | 3.738 | 3.4971 | 3.6175 | 8.5012 | 0 | 1
210 | 1 | 0.081505 | 0.0081374 | 3.6342 | 3.3657 | 3.5 | 8.2249 | 0 | 1
215 | 1 | 0.038124 | 0.11566 | 3.5527 | 3.2917 | 3.4222 | 8.0421 | 0 | 1
220 | 1 | 0.13191 | 0.028305 | 3.4098 | 3.2165 | 3.3132 | 7.786 | 0 | 1
225 | 1 | -0.094419 | 0.02465 | 3.3269 | 3.2357 | 3.2813 | 7.7111 | 0 | 1
230 | 1 | 0.11215 | 0.16858 | 3.2732 | 3.306 | 3.2896 | 7.7306 | 0 | 0.99986
235 | 1 | 0.030276 | -0.04285 | 3.1793 | 3.3105 | 3.2449 | 7.6255 | 0 | 0.99914
240 | 1 | 0.0060168 | 0.24819 | 3.1513 | 3.3035 | 3.2274 | 7.5844 | 0 | 0.99961
245 | 1 | 0.033296 | 0.21182 | 2.966 | 3.0828 | 3.0244 | 7.1074 | 0 | 0.99947